## Slide 1: CHARA AO WFS Design
JDM (+LS,MJI)
2012Sep06
v0.1

## Slide 2: WFS Design Goals
Subaperture geometry matched to Cilas DMs under consideration
Camera frame rate 1 khz
High throughput (few reflections)
Lenslet plane will be conjugate with secondary mirror location
Filters: Bandpass filter (no ADC) and notch filter (laser diode beacon)
Off the shelf components, when possible

## Slide 3: A. Subaperture geometry
Spacing 0.20m
18 spots
Spacing 0.17m
30 spots
Spacing 0.14m
44 spots
Cilas DM
31 Actuator
(were these ellptical?)
Cilas DM
61 Actuator

## Slide 4: Example Placement

## Slide 5: B. Camera Rate
Design goal of 1 Khz frame rate on camera
  Good correction for fast seeing on bright targets
  Avoid saturation for bright targets
  Run camera in photon counting mode
Assuming Andor iXon Ultra 897, 1 khz == 90x90 pixels in fastest readout mode
Issues:
  Need field/pupil stop for ‘cropped sensor mode’
  Shutter?
  Where to put subarray if no shutter (de factor frame transfer ccd)?

## Slide 6: C. High throughput

## Slide 7: Zoom-up.

## Slide 8: D. Location of Lenslet plane
Location of focus
Field stop 5”
(~100micron)
Light from parabola
Collimate beam so aperture is 1.5mm
Achromatic doublet
F=7.5mm, AR coated
Input pupil (secondary) imaged here: 7.2 mm from collimator
Lenslet
goes here

## Slide 9: Example of wavefront quality

## Slide 10: Adding filter upstream improves wavefront

## Slide 11: E. Filters: bandpass + beacon notch
Light from parabla
Filter(s) could go here. Must be interchangeble
Lenslet
goes here
~5mm

## Slide 12: E. Off the shelf components
OKO C-mount hexagonal lenslet
  Spacing 0.3mm, f=18mm
  Used for all calculations
Thorlabs f=7.5mm achromat

## Slide 13: Unresolved Design Issues: Lenslet
Standard OKO lenslet has good features
  Off the shelf
  C-mounted
  AR coated
  FOV of each lenslet is about right (5.2”)
Some issues
  Over sample PSF a bit more than necessary (3.2pixels FWHM), which means more dark noise than necessary when closed loop
  Requires reading out 90x90 for 0.2m aperture or 130x130 (600Hz) for 0.14m apertures (if we want 0.14m apertures we probably require custom lenslet)
  Loses ~15% of light in the edges because circular lenses not hexagonal
  Custom lenslet run will likely be 10000-20000 euros
  If custom lenslet has f<18mm, then requires major modification to the CCD window assembly and mounting (this will happen if we want lower sampling on the 16micron pixel CCD)
Hexagonal or rectangular lenslets?
Is diameter 5.2” enough field of view (+/-2.6”)?
If we don’t get lenses that fill, do we need a chrome mask to eliminate light spillover?

## Slide 14: Unresolved design issues: Bandpass Filters
In front of focus
In converging beam – but surpisingly this seems to correct downstream aberrations in collimating doublet (but depends on which doublet is chosen – not GENERAL result)
  All filters must be matched to not change focus position
Can use off the shelf (larger format possible)
Space to mechanize
*might* compromise acquisition camera sensitivity (throughput/image quality -- TBD)
Behind focus
In collimated beam (good)
Must be smaller (off the shelf?)
Harder to exchange but still possible probably (hard to mechanize)
Very tight fit for two filters (notch+bandpass) but possible probably

## Slide 15: Notes on Camera Modes
Baseline design:
Saturation occurs of V= -0.5mag (no ND filter needed)
1 count per pixel (peak) for V=12.5
Noise: 1.5 dark counts per 5”x5” FOV lenslet (but after we close loop we can focus on 2”x2” region which has only 0.25 dark counts per frame)
We have to run always in 17Mhz multiplication mode and will adjust adjust gain for each object
For FAINT stars we should integrate longer than 1ms to improve SNR since clock-induced counts occur per FRAME not per second for this device
Optimal subarray location is not clear
  minimize blurring of image during readout (assuming we have no shutter)
  Probably need camera to test if offsetting subarray is still fast and leads to “frame transfer ccd” readout
  Should we just get shutter? Will it work in for 1 khz for years?

## Slide 16: Other Questions
How to deal with wandering pupil due to poor coude alignment?
  Do we build in ways to deal with this or rely on good coude alignment?
  Could mount WFS assembly to pivot around hole in center of AQ mirror. This would allow a pupil shift without changing upstream tilts. But probably not so easy to do at the level necessary (ie pivot wander <100microns on large mount with camera, etc).
The hole in the AQ mirror will act as our field stop..
  But if we can’t get the hole small enough we could bond a pinhole. Important to have a field stop here to match field of view of lenslets (~5”)
  Not sure we can buy a AQ mirror with such a small hole
    one vendor I contacted has limit of 1.3mm
Should we allow for off-axis guiding?
  I think yes, which means large angular throw using encoders for either fold mirror or parabola

## Slide 17: Next steps
Is basic design architecture ok?
  Problems? Concerns? Risks?
Decide what options we want to include for trade study for presentation to critical design review
Do sensitivity calculation
  simulation of wavefronts, camera detection, wavefront reconstruction, and resulting performance (TT only, TT+AO)
Get quotes for custom items
Do test on lab camera of critical issues
  Confirm readout speeds, noise properties, latency
  Check mechanical constraints on window
Proceed to design study of acquisition camera/optics
Proceed to design for the beacon/stimulus/alignment system to see how it impacts design of WFS
Check mechanical layouts
  Clearance with dicroic changer, other beams, size of table, placeholder for fiber injection module